PROJECT TITAN

Educational Risk Analysis (FMEA)

Educational Engineering Project

| Document Number | TITAN-RSK-001 |
| --- | --- |
| Revision | A |
| Status | Released - Educational Baseline |
| Author | Michael Lozada-Longog |
| Affiliation | Biomedical Engineering, Hawaii Pacific University |
| Date | July 2026 |
| Classification | Educational Engineering Project |

IMPORTANT SCOPE NOTICE

Project TITAN is an educational design and analysis project. It is not a clinically validated device, regulatory submission, or product cleared for human use.

# Document Control and Revision History

| Revision | Date | Description | Prepared By | Status |
| --- | --- | --- | --- | --- |
| A | July 2026 | Initial educational design FMEA for mechanical, manufacturing, assembly, and use-related hazards. | Michael Lozada-Longog | Released |

Source baseline: Project TITAN Engineering Design Report (July 2026), CAD assembly, completed engineering drawings, and Fusion 360 static FEA results documented by the author.

# 1. Purpose and Method

This document applies a simplified Failure Modes and Effects Analysis to Project TITAN. It is intended to demonstrate risk-based engineering thinking. It is not an ISO 14971-compliant risk-management file, clinical hazard analysis, or regulatory submission.

Scales: Severity (S), Occurrence (O), and Detection (D) are rated 1–10. Risk Priority Number (RPN) = S × O × D. Ratings are preliminary engineering estimates and must be updated with prototype and test data.

# 2. Rating Guidance

| Rating | Severity | Occurrence | Detection |
| --- | --- | --- | --- |
| 1–2 | Negligible effect | Remote | Almost certain detection |
| 3–4 | Minor functional degradation | Low | High likelihood of detection |
| 5–6 | Loss of intended function or repair required | Moderate | Moderate likelihood of detection |
| 7–8 | Major structural/function loss; potential injury in a future wearable context | High | Low likelihood of detection |
| 9–10 | Catastrophic failure or severe injury potential | Very high | Unlikely to be detected before effect |

# 3. Preliminary Design FMEA

| ID | Failure Mode | Cause | Effect | S | O | D | RPN | Recommended Control |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FM-01 | Bearing seizure or binding | Contamination, misalignment, inadequate clearance | Loss of smooth rotation; increased load into structure | 7 | 4 | 5 | 140 | Use sealed COTS bearings; define fits; verify alignment and rotational torque |
| FM-02 | Bearing loosens in plate bore | Insufficient retention or incorrect fit | Loss of alignment and local wear | 6 | 4 | 6 | 144 | Add shoulder/retainer; specify fit and inspection method |
| FM-03 | Side plate yields or cracks | Overload, stress concentration, cyclic fatigue | Permanent deformation or structural separation | 9 | 3 | 5 | 135 | Maintain SF target; add fatigue study; inspect high-stress regions |
| FM-04 | Main shaft bends | Excess radial load or inadequate shaft material | Misalignment, bearing damage, loss of motion | 8 | 3 | 5 | 120 | Hand-calculate bending; reconsider steel shaft; test deflection |
| FM-05 | Main shaft surface wears | Aluminum shaft against bearing/interface; contamination | Clearance growth and poor alignment | 6 | 5 | 6 | 180 | Use hardened steel/stainless shaft or hardened sleeve; define lubrication |
| FM-06 | Spacer wear or deformation | Incorrect material, edge loading, poor fit | Axial play and component rubbing | 5 | 4 | 5 | 100 | Specify material and flatness; verify stack-up |
| FM-07 | Fastener loosening | Vibration, insufficient torque, no locking feature | Joint motion and progressive damage | 8 | 4 | 5 | 160 | Define fastener grade, torque, locking method, inspection interval |
| FM-08 | Fastener pull-through or hole elongation | Thin section, high bearing stress, repeated motion | Loss of rigid frame connection | 8 | 3 | 6 | 144 | Check bearing stress; add washers/bushings; inspect holes |
| FM-09 | Upper/lower frame bends excessively | High moment, insufficient stiffness | Poor alignment and reduced function | 7 | 3 | 5 | 105 | Verify displacement; increase section or add ribs if needed |
| FM-10 | Interference during rotation | Inadequate clearance or assembly error | Binding, sudden stop, part damage | 6 | 4 | 3 | 72 | CAD interference review; physical motion test; clearance requirement |
| FM-11 | Incorrect assembly order | Ambiguous instructions or mirrored parts | Misalignment, interference, incomplete engagement | 5 | 4 | 3 | 60 | Number parts; create assembly instructions and inspection checklist |
| FM-12 | Incorrect material installed | Procurement or labeling error | Unexpected strength, corrosion, or wear behavior | 7 | 2 | 6 | 84 | Material certs; part marking; incoming inspection |
| FM-13 | Sharp edges or burrs | Incomplete deburring | Cuts, snagging, stress risers | 6 | 4 | 2 | 48 | Deburr and edge-break requirement; visual inspection |
| FM-14 | Corrosion or sweat exposure damage | Humidity, salts, dissimilar-metal contact | Surface degradation, seizure, reduced life | 6 | 4 | 6 | 144 | Anodize/coat; isolate dissimilar metals; environmental test |
| FM-15 | FEA result misinterpreted | Idealized contacts, outdated mesh, incorrect constraints | Unsafe design decision based on invalid evidence | 9 | 3 | 7 | 189 | Document assumptions; mesh check; peer review; correlate with test |
| FM-16 | Load applied asymmetrically | Incorrect use or uneven interface | Unexpected torsion and one-sided plate stress | 8 | 4 | 6 | 192 | Add off-axis load case and torsional analysis |
| FM-17 | Range of motion exceeds safe mechanical limit | No validated hard stop in baseline | Hyperextension or component collision in future wearable concept | 9 | 4 | 7 | 252 | Design adjustable hard stop; verify angular limits before human-use consideration |
| FM-18 | Device used on a person before validation | Scope misunderstanding | Potential injury | 10 | 2 | 8 | 160 | Prominent educational-use labeling; prohibit human testing without formal review |

# 4. Risk Prioritization

| Priority | ID | Failure Mode | Initial RPN | Disposition |
| --- | --- | --- | --- | --- |
| 1 | FM-17 | Range of motion exceeds safe mechanical limit | 252 | Action required before physical/wearable V2 |
| 2 | FM-16 | Load applied asymmetrically | 192 | Action required before physical/wearable V2 |
| 3 | FM-15 | FEA result misinterpreted | 189 | Action required before physical/wearable V2 |
| 4 | FM-05 | Main shaft surface wears | 180 | Action required before physical/wearable V2 |
| 5 | FM-18 | Device used on a person before validation | 160 | Track and verify |
| 6 | FM-07 | Fastener loosening | 160 | Track and verify |
| 7 | FM-14 | Corrosion or sweat exposure damage | 144 | Track and verify |
| 8 | FM-08 | Fastener pull-through or hole elongation | 144 | Track and verify |

# 5. Required Risk Controls Before Prototype

Revisit the main shaft material and bearing interface.

Define bearing retention and fit tolerances.

Select and document fasteners, torque, and locking methods.

Add an off-axis/torsional load case and preliminary fatigue analysis.

Create assembly, inspection, and labeling instructions.

Add prominent “not for human use” labeling to the prototype and technical package.

Plan physical static-load and deflection tests before any wearable evaluation.

# 6. Verification of Risk Controls

| Control | Verification Method | Acceptance Evidence | Status |
| --- | --- | --- | --- |
| Minimum structural safety factor | FEA and later physical load test | FEA SF ≥2.0; test criterion TBD | FEA complete; test planned |
| Bearing alignment | CAD inspection and rotational torque test | No binding; torque limit TBD | CAD complete |
| Bearing retention | Drawing review and pull-out inspection | Retention design and test value TBD | Open |
| Fastener security | Torque spec and cyclic inspection | No loosening after defined cycles | Open |
| Edge safety | Visual and tactile inspection | No burrs or sharp edges | Future fabrication |
| Scope control | Label/document review | Educational-use notice present | Documented |

# 7. Residual Risk Statement

Residual risk cannot be formally accepted because no physical prototype, fatigue test, misuse evaluation, or human-factors assessment has been completed. The current FMEA supports design planning only. Project TITAN must remain a non-clinical educational device until future verification evidence is generated and reviewed.

# 8. FMEA Maintenance

Update this FMEA whenever the geometry, material, fastener strategy, load case, bearing retention, prototype process, or intended use changes. New test data should replace preliminary occurrence and detection ratings.

# References

TITAN-PRD-001, Product Requirements Document, Revision B.

TITAN-SYS-001, System Architecture Document, Revision A.

Project TITAN Engineering Design Report, July 2026.