PROJECT TITAN

System Architecture Document

Educational Engineering Project

| Document Number | TITAN-SYS-001 |
| --- | --- |
| Revision | A |
| Status | Released - Educational Baseline |
| Author | Michael Lozada-Longog |
| Affiliation | Biomedical Engineering, Hawaii Pacific University |
| Date | July 2026 |
| Classification | Educational Engineering Project |

IMPORTANT SCOPE NOTICE

Project TITAN is an educational design and analysis project. It is not a clinically validated device, regulatory submission, or product cleared for human use.

# Document Control and Revision History

| Revision | Date | Description | Prepared By | Status |
| --- | --- | --- | --- | --- |
| A | July 2026 | Initial system architecture baseline for the Project TITAN mechanical knee-support assembly. | Michael Lozada-Longog | Released |

Source baseline: Project TITAN Engineering Design Report (July 2026), CAD assembly, completed engineering drawings, and Fusion 360 static FEA results documented by the author.

# 1. Purpose and Scope

This document defines the mechanical architecture, functional decomposition, component hierarchy, interfaces, load path, assembly sequence, and analytical boundaries of Project TITAN. It translates the requirements in TITAN-PRD-001 into a physical system definition suitable for CAD, FEA, manufacturing planning, and future prototype development.

Architecture Baseline The current baseline is a nine-component structural assembly: main shaft, left and right side plates, upper and lower frames, left and right spacers, and two commercial bearings. No powered actuator, sensor, strap, human-interface pad, or clinically intended component is included in the released baseline.

# 2. System Mission

The system mission is to provide a modular mechanical platform that supports a bearing-guided rotational interface while maintaining structural alignment under a defined static load case. The architecture is intentionally simple so that structural behavior, manufacturing documentation, and design decisions can be evaluated clearly.

# 3. Functional Decomposition

| Function ID | Primary Function | Allocated Elements | Verification |
| --- | --- | --- | --- |
| F-01 | Carry external structural loads | Side plates, upper frame, lower frame | Static FEA and drawing review |
| F-02 | Provide rotational support | Main shaft and two bearings | Assembly inspection |
| F-03 | Maintain axial spacing and alignment | Left and right spacers | CAD dimensional review |
| F-04 | Connect load-carrying members | Upper/lower frames and side plates | Assembly and interface review |
| F-05 | Enable component replacement | Modular component architecture | BOM and assembly review |
| F-06 | Provide analyzable geometry | Simplified parametric parts | FEA preprocessing and solve completion |

# 4. Physical Architecture

The released design uses a dual-plate arrangement surrounding a central shaft and bearing stack. Upper and lower frame members establish the principal longitudinal structure. Spacers preserve axial separation and reduce direct side loading between adjacent components.

System hierarchy:

Project TITAN Assembly

Structural Subsystem: left plate, right plate, upper frame, lower frame

Rotational Subsystem: main shaft, left bearing, right bearing

Alignment Subsystem: left spacer, right spacer

Documentation Subsystem: CAD, drawings, BOM, FEA, engineering report

# 5. Component Architecture and Responsibilities

| Part No. | Component | Type | Primary Responsibility | Baseline Material |
| --- | --- | --- | --- | --- |
| TITAN-001 | Main Shaft | Custom | Defines the central rotational axis and transfers load into the bearing interfaces | 6061-T6 aluminum in baseline documentation |
| TITAN-002 | Left Side Plate | Custom | Carries structural load and supports one bearing seat | 6061-T6 aluminum |
| TITAN-003 | Right Side Plate | Custom | Provides opposing support and completes the dual-plate load path | 6061-T6 aluminum |
| TITAN-004 | Upper Frame | Custom | Provides the upper longitudinal structural connection and strap-slot concept geometry | 6061-T6 aluminum |
| TITAN-005 | Lower Frame | Custom | Provides the lower longitudinal structural connection and strap-slot concept geometry | 6061-T6 aluminum |
| TITAN-006 | Left Spacer | Custom | Maintains left-side axial clearance and alignment | 6061-T6 aluminum |
| TITAN-007 | Right Spacer | Custom | Maintains right-side axial clearance and alignment | 6061-T6 aluminum |
| TITAN-008 | Left Bearing | COTS concept | Provides rolling support at the left interface | 52100 bearing steel |
| TITAN-009 | Right Bearing | COTS concept | Provides rolling support at the right interface | 52100 bearing steel |

# 6. Mechanical Interfaces

| Interface ID | Elements | Interface Type | Design Intent | Key Concern |
| --- | --- | --- | --- | --- |
| IF-01 | Main shaft ↔ bearing inner surfaces | Cylindrical load-transfer interface | Transfer radial bearing load while preserving axis alignment | Fit and contact definition |
| IF-02 | Bearing outer surfaces ↔ side-plate bores | Bearing-seat interface | Transfer load into structural plates | Press-fit/retention not yet production-defined |
| IF-03 | Spacers ↔ shaft/adjacent components | Axial spacing interface | Prevent component rubbing and control stack width | Tolerance stack-up |
| IF-04 | Upper frame ↔ side plates | Rigid structural connection | Transfer upper-frame load into plate system | Fastener details remain conceptual |
| IF-05 | Lower frame ↔ side plates | Rigid structural connection | Transfer lower-frame load into plate system | Fastener details remain conceptual |

# 7. Load Path

The validated educational FEA load case applied bearing loads to the inner cylindrical bearing surfaces. The intended idealized load path is:

Load Path Applied bearing load → bearing inner surface → bearing body/contact definition → side-plate bearing seat → side plates → upper/lower frame support regions → fixed structural boundary.

The final refined-mesh study reported a maximum von Mises stress of 54.95 MPa, minimum safety factor of 2.377, and maximum displacement of 0.004 in (approximately 0.10 mm). These values are accepted only for the documented idealized boundary conditions and do not constitute clinical validation.

# 8. Assembly Sequence

Prepare and inspect custom components and commercial bearings.

Seat the left and right bearings into their corresponding plate bores.

Position the main shaft through the bearing stack.

Install left and right spacers to establish the intended axial stack.

Align the second side plate with the shaft/bearing stack.

Attach the upper frame to the side-plate structure.

Attach the lower frame to the side-plate structure.

Perform interference, alignment, and free-rotation checks.

Record the as-built configuration and revision.

# 9. Analysis Architecture

The project used two levels of structural analysis: an isolated left-side-plate study for setup validation and a full-assembly static study for system-level evaluation. The full-assembly study required explicit contacts, constraints, loads, mesh generation, and cleanup of duplicate or unconstrained bodies.

| Analysis Level | Purpose | Inputs | Outputs | Limitation |
| --- | --- | --- | --- | --- |
| Component | Validate plate behavior and load application | 6061-T6, fixed mounting holes, bearing load | Stress, displacement, safety factor | Does not capture assembly interactions |
| Full assembly | Evaluate integrated load path | All retained bodies, contacts, fixed supports, bearing loads | System stress, displacement, safety factor | Idealized contacts and supports |

# 10. Architecture Constraints and Open Items

The current design does not include a production-defined bearing retention shoulder, snap ring, adhesive, or interference fit.

Fastener size, grade, thread engagement, torque, and locking method are not yet released.

Human-contact surfaces, straps, padding, and anthropometric adjustability are outside the current baseline.

The main shaft material is documented as aluminum in the current drawing package; a higher-strength shaft material remains a future trade-study item.

Dynamic, fatigue, wear, sweat exposure, corrosion, and misuse conditions have not been experimentally validated.

# 11. Future Architecture Evolution

| Version | Architecture Goal | Planned Additions |
| --- | --- | --- |
| V2 | Functional mechanical prototype | Production-like fits, retention, fasteners, improved geometry |
| V3 | Instrumented biomedical concept | IMU, angle sensing, load sensing, data acquisition |
| V4 | Connected rehabilitation platform | Bluetooth, dashboard, data logging |
| V5 | Evaluation platform | Repeatability, usability, benchtop validation, human-factors planning |

# 12. Conclusion

The Project TITAN architecture establishes a clear relationship between requirements, physical components, interfaces, and verification evidence. The current system is a mechanically focused educational baseline. Its principal strengths are modularity, analyzability, manufacturability awareness, and traceable documentation; its principal maturity gaps are production hardware definition, physical prototyping, and experimental validation.

# References

Project TITAN Engineering Design Report, Michael Lozada-Longog, July 2026.

TITAN-PRD-001, Product Requirements Document, Revision B.

Project TITAN CAD assembly and Fusion 360 simulation outputs, July 2026.